## initial
clause | type | display.text | comments
do section survey
goto _finalize |  |  | skips the finalize screen where the user chooses to save as incomplete or finalized and instead saves as finalized

## survey
clause | condition | type | values_list | name | display.prompt.text | display.prompt.text.es | display.prompt.text.fr | display.hint.text | display.hint.text.es | display.hint.text.fr | required | calculation | constraint | display.constraint_message.text | display.constraint_message.text.fr | display.title.text
begin screen
if | data('refrigerator_id') === null ¦¦ data('refrigerator_id') === undefined
|  | string |  | refrigerator_id | Refrigerator ID: | ID de Frigorífico | ID de réfrigérateur | Enter the ID of the refrigerator | Por favor entre el ID del frigorífico |  | TRUE
end if
|  | note |  |  | <h2 style="font-family: Arial, sans-serif; color: #2c3e50;">CCE Identification</h2> <table style="border-collapse: collapse; width: 100%; font-family: Arial, sans-serif;"> <tr style="background-color: #f2f2f2;"> <th style="border: 1px solid #ccc; padding: 10px; text-align: left;">Field</th> <th style="border: 1px solid #ccc; padding: 10px; text-align: left;">Value</th> </tr> <tr> <td style="border: 1px solid #ccc; padding: 10px;"><b>Refrigerator ID</b></td> <td style="border: 1px solid #ccc; padding: 10px;">{{data.refrigerator_id}}</td> </tr> <tr> <td style="border: 1px solid #ccc; padding: 10px;"><b>Brand</b></td> <td style="border: 1px solid #ccc; padding: 10px;">{{data.manufacturer}}</td> </tr> <tr> <td style="border: 1px solid #ccc; padding: 10px;"><b>Model</b></td> <td style="border: 1px solid #ccc; padding: 10px;">{{data.model_id}}</td> </tr> <tr> <td style="border: 1px solid #ccc; padding: 10px;"><b>Health Facility</b></td> <td style="border: 1px solid #ccc; padding: 10px;">{{data.facility_name}}</td> </tr> </table>
|  | select_one_dropdown | onsiteremote | onsite_remote | Are you onsite or remote?
end screen
|  | note |  |  | <div style="padding: 15px; border-left: 5px solid #e67e22; background-color: #fef6f0; font-family: Arial, sans-serif;"> <h3 style="margin-top: 0; color: #2c3e50;">Section A</h3> <p style="font-size: 16px; color: #34495e;"> This section can be completed by the Surveillance Officer either remotely or on-site. </p> </div>
begin screen |  |  |  |  | A
|  | note |  | title_section_a | <div style="padding: 15px; border-left: 5px solid #e67e22; background-color: #fef6f0; font-family: Arial, sans-serif;"> <h3 style="margin-top: 0; color: #2c3e50;">Section A</h3></div>
|  | select_one_with_other | refrigerator_place | refrigerator_place | 1. Where is the FT2 (or FT2 sensor) placed inside the refrigerator? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Ask where the FT2/sensor is usually placed. Confirm if it stays in one place or is moved (e.g., to access vaccines).\n\nIf onsite: visually verify the placement and note any differences.\n\nIf the sensor is regularly moved, select 'Other' and explain.`)">[?]</span> | [7 empty] | data('refrigerator_place') != ' ' | Other field is empty. Please specify.
|  | select_one | yesnou | temperature_alarm_caused | 2. Do any temperature alarms appear to have been caused by the removal of FT2 (or FT2 sensor) from refrigerator? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Review FT2 data for gaps in readings or signs the sensor was removed (e.g., abrupt/extended ambient readings).\n\nIf such gaps match the timing of alarms, select 'Yes'.\nIf all alarms occurred while the sensor was in the refrigerator, select 'No'.\nIf unclear or inconsistent, select 'Unknown'.`)">[?]</span>
|  | select_one_with_other | time | refrigerator_cleaned | 3. How often is the refrigerator cleaned? | [7 empty] | data('refrigerator_cleaned') != ' ' | Other field is empty. Please specify.
|  | select_one_with_other | time | refrigerator_drained | 4. How often is the refrigerator drained? | [7 empty] | data('refrigerator_drained') != ' ' | Other field is empty. Please specify.
|  | select_one | yesnou | refrigerator_exposed | 5. Is the refrigerator in a location exposed to direct sunlight, rainwater, or other harsh conditions at any times or seasons? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Ask the user and visually check for roof leaks that could expose equipment to rain.\n\nIf remote: Ask the user.`)">[?]</span>
|  | select_one | yesno | refrigerator_ventilation_space | 6. Does the refrigerator have at least 2 cm of space for ventilation on all sides? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Use a ruler, if available, or estimate space on all sides of the refrigerator, including the bottom. Ensure electronics vents are not obstructed.\n\nIf remote: Ask the user to estimate or measure space on all sides and confirm that vents are not blocked.`)">[?]</span>
|  | decimal |  | reading_temperature_display | 7. What is the current reading on the external temperature display in ˚C (if applicable)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Record the temperature reading currently shown on the external display.\n\nIf remote: Ask the user to read and report the current temperature from the external display.`)">[?]</span>
|  | select_one | yesno | water_collected_refrigerator | 8. Is there currently any water collected in the bottom of the refrigerator? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Look inside the refrigerator for any water pooling at the bottom of the vaccine compartment.\n\nIf remote: Instruct the user to do the same and report back.`)">[?]</span>
if | selected(data('water_collected_refrigerator'), 'Yes')
|  | decimal |  | water_height | 8.1. What is the height of the water in cm? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Use a ruler, if available, or estimate the height of the water collected at the bottom of the vaccine compartment.\n\nIf remote: Instruct the user to do the same and report back.`)">[?]</span>
end if
|  | select_one | yesno | store_vaccines | 9. Are baskets or shelves used to store vaccines? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Look inside the vaccine compartment and observe how vaccines are stored. If all vaccines are on a shelf or in a suspended basket, answer Yes. Otherwise, answer No.\n\nIf remote: Instruct the user to check inside the vaccine compartment and report how vaccines are stored. If all are on a shelf or in a suspended basket, answer Yes. Otherwise, answer No.`)">[?]</span>
|  | select_one | yesno | gasket_damage | 10. Does the gasket have any visible damage, such as rips, deadhesion, or mold? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Check the gasket for any visible signs of damage.\n\nIf remote: Instruct the user to check the gasket for any visible signs of damage.`)">[?]</span>
if | selected(data('gasket_damage'), 'Yes')
|  | string |  | gasket_observation | 10.1. Describe observation: <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Enter details about gasket damage observed.\n\nIf remote: Ask the user for details about gasket damage observed.`)">[?]</span> | [11 empty] | hint
|  | image |  | gasket_observation_image | 10.2. Upload Image of the damage
end if
|  | select_one | yesno | corrosion_refrigerator | 11. Is there any corrosion visible on the refrigerator exterior, such as on hinges, hardware, or vents? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;"onclick="alert(`If onsite: Check the refrigerator exterior for any signs of corrosion, particularly on the hinges, hardware, and vents.\n\nIf remote: Instruct the user to check the refrigerator exterior for any signs of corrosion, particularly on the hinges, hardware, and vents.`)">[?]</span>
if | selected(data('corrosion_refrigerator'), 'Yes')
|  | string |  | corrosion_obervation | 11.1. Describe observation: <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Enter details about corrosion observed on the refrigerator's exterior (e.g., hinges, hardware, vents).\n\nIf remote: Ask the user to describe any corrosion observed on the refrigerator’s exterior.`)">[?]</span>
|  | image |  | corrosion_obervation_image | 11.2. Upload Image of the corrosion
end if
|  | select_one | yesno | condensation_drain_problems | 12. Is the condensation drain leaking, blocked, or missing the plug? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Check the vaccine compartment drain for leaks outside the refrigerator, blockages inside, and confirm the drain plug is properly in place.\n\nIf remote: Instruct the user to check the drain for leaks or blockages and confirm the plug is securely in place.`)">[?]</span>
if | selected(data('condensation_drain_problems'), 'Yes')
|  | string |  | condensation_drain_observation | 12.1. Describe observation: <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Enter details about any visible issues with the condensation drain.\n\nIf remote: Ask the user to describe any problems they've noticed with the condensation drain.`)">[?]</span>
|  | image |  | condensation_drain_observation_image | 12.2. Upload Image of the condensation drain leaking, blocked or missing the plug
end if
|  | select_one | yesno | power_cable_damage | 13. Does the power cable and/or plug show any signs of damage from rodents, strain, or other causes? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;"onclick="alert(`If onsite: Check the refrigerator power cable and plug for damage from rodents, strain (e.g., stretched cable), or other visible issues.\n\nIf remote: Instruct the user to check the refrigerator power cable and plug for any such damage or strain.`)">[?]</span>
if | selected(data('power_cable_damage'), 'Yes')
|  | string |  | power_cable_damage_observation | 13.1. Describe observation: <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Enter details about any damage observed to the power cable or plug (e.g., rodent bites, fraying, stretched cable).\n\nIf remote: Ask the user to describe any visible damage to the power cable or plug and enter their response.`)">[?]</span>
|  | image |  | power_cable_damage_image | 13.2. Upload image of the damaged power cable and/or plug
end if
|  | select_one | yesnou | refrigerator_noise_smell | 14. Has the refrigerator been making any loud noises or producing abnormal smells during the past 30 days? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Ask user`)">[?]</span>
if | selected(data('refrigerator_noise_smell'), 'Yes')
|  | string |  | refrigerator_noise_smell_observation | 14.1. Describe observation: <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Enter any unusual noises or smells observed during the visit.\n\nIf remote: Ask the user if they have noticed any unusual noises or smells from the refrigerator over the past 30 days, and record their response.`)">[?]</span>
|  |  |  | refrigerator_noise_smell_image | 14.2. Upload image to support observation
end if
if | data('power_source') === 'solar'
|  | select_one_with_other | solar_array_location | solar_array_location | 15. Where is the solar array located? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Ask the user and visually verify the location of the solar array.\n\nIf remote: Ask the user about the location of the solar array.`)">[?]</span> | [7 empty] | data('solar_array_location') != ' ' | Other field is empty. Please specify.
|  | select_one | yesno | repair_solar_system | 16. In the past 3 months (or since the last PMM check), has anyone been here to repair or work on the solar system, or has repair or maintenance been requested for any reason? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Ask the user and verify with repair logs if available.\n\nIf remote: Ask the user.`)">[?]</span>
if | selected(data('repair_solar_system'), 'Yes')
|  | string |  | repair_why_what_actions | 16.1. Why? What system/area? What action was taken? What parts were replaced (if any)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: Ask the user and verify with repair logs if available.\n\nIf remote: Ask the user.`)">[?]</span>
end if
|  | select_one_with_other | time | solar_panels_cleaned | 17. How often are the solar panels cleaned? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Ask the user.`)">[?]</span> | [7 empty] | data('solar_panels_cleaned') != ' ' | Other field is empty. Please specify.
end if
|  | select_one | yesno_section_a | failure_cause_known | 18. Has the cause of the failure been identified? | [5 empty] | TRUE
end screen
if | selected(data('failure_cause_known'), 'No')
|  | note |  |  | <div style="padding: 15px; border-left: 5px solid #e67e22; background-color: #fef6f0; font-family: Arial, sans-serif;"> <h3 style="margin-top: 0; color: #2c3e50;">Section B</h3> <p style="font-size: 16px; color: #34495e;"> To complete this section the Surveillance Officer must work with a technician or informed staff who is on-site and able to investigate the appliance directly. </p></div>
begin screen
|  | note |  | title_section_b | <div style="padding: 15px; border-left: 5px solid #e67e22; background-color: #fef6f0; font-family: Arial, sans-serif;"> <h3 style="margin-top: 0; color: #2c3e50;">Section B</h3> </div>
|  | select_one_with_other | main_component | main_component_first_reported_category | 18. What is the type of main component associated with the first reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: If at least one build quality issue was reported in the most recent sentinel survey, investigate the first issue noted. Ask follow-up questions and visually verify the component to ensure accuracy (e.g., distinguish between corroded pipes vs. filter dryer).\n\nIf remote: Section B cannot be completed remotely, ask user who is on-site.`)">[?]</span>
if | selected(data('main_component_first_reported_category'), 'cooling')
|  | select_one_with_other | cooling_system | main_component_first_cooling | 18.1. What is the main component associated with the first reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: If at least one build quality issue was reported in the most recent sentinel survey, investigate the first issue noted. Ask follow-up questions and visually verify the component to ensure accuracy (e.g., distinguish between corroded pipes vs. filter dryer).\n\nIf remote: Section B cannot be completed remotely, ask user who is on-site.`)">[?]</span>
end if
if | selected(data('main_component_first_reported_category'), 'electrical')
|  | select_one_with_other | electrical_system | main_component_first_electrical | 18.1. What is the main component associated with the first reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: If at least one build quality issue was reported in the most recent sentinel survey, investigate the first issue noted. Ask follow-up questions and visually verify the component to ensure accuracy (e.g., distinguish between corroded pipes vs. filter dryer).\n\nIf remote: Section B cannot be completed remotely, ask user who is on-site.`)">[?]</span>
end if
if | selected(data('main_component_first_reported_category'), 'electrical_solar')
|  | select_one_with_other | electrical_solar_system | main_component_first_solar | 18.1. What is the main component associated with the first reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: If at least one build quality issue was reported in the most recent sentinel survey, investigate the first issue noted. Ask follow-up questions and visually verify the component to ensure accuracy (e.g., distinguish between corroded pipes vs. filter dryer).\n\nIf remote: Section B cannot be completed remotely, ask user who is on-site.`)">[?]</span>
end if
if | selected(data('main_component_first_reported_category'), 'structural')
|  | select_one_with_other | structural_components | main_component_first_structural | 18.1. What is the main component associated with the first reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If onsite: If at least one build quality issue was reported in the most recent sentinel survey, investigate the first issue noted. Ask follow-up questions and visually verify the component to ensure accuracy (e.g., distinguish between corroded pipes vs. filter dryer).\n\nIf remote: Section B cannot be completed remotely, ask user who is on-site.`)">[?]</span>
end if
|  | select_one_with_other | main_component | main_component_second_reported_category | 19. What is the type of main component associated with the second reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least two build quality issues were reported in the most recent sentinel survey:\n\nOnsite: Investigate the second issue. Read the sentinel survey, ask follow-up questions, and visually verify the component. Ensure accurate identification (e.g., confirm if corrosion is on the filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite to assist.`)">[?]</span>
if | selected(data('main_component_second_reported_category'), 'cooling')
|  | select_one_with_other | cooling_system | main_component_second_cooling | 19.1. What is the main component associated with the second reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least two build quality issues were reported in the most recent sentinel survey:\n\nOnsite: Investigate the second issue. Read the sentinel survey, ask follow-up questions, and visually verify the component. Ensure accurate identification (e.g., confirm if corrosion is on the filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite to assist.`)">[?]</span>
end if
if | selected(data('main_component_second_reported_category'), 'electrical')
|  | select_one_with_other | electrical_system | main_component_second_electrical | 19.1. What is the main component associated with the second reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least two build quality issues were reported in the most recent sentinel survey:\n\nOnsite: Investigate the second issue. Read the sentinel survey, ask follow-up questions, and visually verify the component. Ensure accurate identification (e.g., confirm if corrosion is on the filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite to assist.`)">[?]</span>
end if
if | selected(data('main_component_second_reported_category'), 'electrical_solar')
|  | select_one_with_other | electrical_solar_system | main_component_second_solar | 19.1. What is the main component associated with the second reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least two build quality issues were reported in the most recent sentinel survey:\n\nOnsite: Investigate the second issue. Read the sentinel survey, ask follow-up questions, and visually verify the component. Ensure accurate identification (e.g., confirm if corrosion is on the filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite to assist.`)">[?]</span>
end if
if | selected(data('main_component_second_reported_category'), 'structural')
|  | select_one_with_other | structural_components | main_component_second_structural | 19.1. What is the main component associated with the second reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least two build quality issues were reported in the most recent sentinel survey:\n\nOnsite: Investigate the second issue. Read the sentinel survey, ask follow-up questions, and visually verify the component. Ensure accurate identification (e.g., confirm if corrosion is on the filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite to assist.`)">[?]</span>
end if
|  | select_one_with_other | main_component | main_component_third_reported_category | 20. What is the type of main component associated with the third reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least three build quality issues were reported:\n\nOnsite: Review the third issue in the sentinel survey. Ask detailed follow-up questions to identify the main component and visually verify it (e.g., corrosion on filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
if | selected(data('main_component_third_reported_category'), 'cooling')
|  | select_one_with_other | cooling_system | main_component_third_cooling | 20.1. What is the main component associated with the third reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least three build quality issues were reported:\n\nOnsite: Review the third issue in the sentinel survey. Ask detailed follow-up questions to identify the main component and visually verify it (e.g., corrosion on filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
end if
if | selected(data('main_component_third_reported_category'), 'electrical')
|  | select_one_with_other | electrical_system | main_component_third_electrical | 20.1. What is the main component associated with the third reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least three build quality issues were reported:\n\nOnsite: Review the third issue in the sentinel survey. Ask detailed follow-up questions to identify the main component and visually verify it (e.g., corrosion on filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
end if
if | selected(data('main_component_third_reported_category'), 'electrical_solar')
|  | select_one_with_other | electrical_solar_system | main_component_third_solar | 20.1. What is the main component associated with the third reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least three build quality issues were reported:\n\nOnsite: Review the third issue in the sentinel survey. Ask detailed follow-up questions to identify the main component and visually verify it (e.g., corrosion on filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
end if
if | selected(data('main_component_third_reported_category'), 'structural')
|  | select_one_with_other | structural_components | main_component_third_structural | 20.1. What is the main component associated with the third reported build quality issue in the sentinel survey? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`If at least three build quality issues were reported:\n\nOnsite: Review the third issue in the sentinel survey. Ask detailed follow-up questions to identify the main component and visually verify it (e.g., corrosion on filter dryer vs. pipes).\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
end if
|  | select_one | yesno | door_hinges_loose | 21. Are the door hinges loose or misaligned? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\nCheck if the hinges are loose and if the door moves due to them.\nUse a ruler to measure the distance between the door edge and cabinet at two points — near the hinges and on the free side.\nThe difference is the door alignment gap:\n- If 0: door is aligned\n- If nonzero: record the gap in the next question\n\nRemote:\nSection B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
if | selected(data('door_hinges_loose'), 'Yes')
|  | decimal |  | door_alignment_gap | 21.1. Door alignment gap (in mm)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\nSee instructions for question 28. \n\nRemote:\nSection B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
end if
|  | select_one | yes_no_unable | access_cover_tempering | 22. Does the access cover show signs of tampering? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\nCheck the access cover for any signs of tampering or unauthorized access.\n\nRemote:\nSection B cannot be completed remotely. Ask a user who is onsite for assistance.`)">[?]</span>
|  | select_one | yes_no_unable | power_led_on_refrigerator | 23. Does the power LED on the refrigerator turn on when the equipment is receiving power (if applicable)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\nCheck the power LED while the equipment is receiving power.\n- For AC systems: verify when utility power is on.\n- For DC systems: check during daylight hours when the solar system is active.\n\nRemote:\nSection B cannot be completed remotely. Please ask a user who is onsite.`)">[?]</span>
|  | select_one | iceliner_intact | ice_liner_intact | 24. Is the ice liner intact (i.e. not missing any water packs and appears to be full) and in the correct place? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\n- Check if an ice liner (coolant panel, water packs, etc.) surrounds the storage area.\n- Refer to the user manual or labels.\n- For B Medical fridges: panels with 6 water packs are common.\n • If packs are missing, empty, or displaced — liner is not intact.\n- If there’s a reservoir with a fill line, ensure coolant reaches the line.\n- If access requires tools or is sealed — select ‘Unknown’.\n\nRemote:\nSection B cannot be completed remotely. Ask a user who is onsite to check.`)">[?]</span>
|  | string |  | thermostat_setpoint | 25. What is the thermostat setpoint? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\n- Check the thermostat set point.\n- If in the electronics compartment, open it using tools (if experienced).\n- Some models require keypad sequences to view the setting.\n- For dial thermostats (e.g., 1–7), note dial position if temperature isn’t shown.\n- If unsure or inaccessible, record as 'Unknown'.\n\nRemote:\nCannot be completed remotely. Ask someone onsite.`)">[?]</span>
|  | select_one | yesno | signs_problem_electronic_compartment | 26. Are there any signs of corrosion, burns, shorts, or heavily accumulated dirt in the electronics compartment? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\n- Open electronics compartment with tools.\n- Check for corrosion, burns, or dirt.\n- Focus on: piping, filter dryer, capillaries, fan, compressor.\n- Record any damage.\n\nRemote:\nCannot be completed remotely. Ask a user onsite.`)">[?]</span>
if | selected(data('signs_problem_electronic_compartment'), 'Yes')
|  | string |  | signs_problem_electronic_compartment_detail | 26.1. Describe the problem, location/component(s), and severity. <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\nEnter detailed observations of the problem (e.g., location, damage, malfunction).\n\nRemote:\nSection B cannot be completed remotely. Please request input from an onsite user.`)">[?]</span>
|  | image |  | signs_problem_electronic_compartment_image | 26.2. Upload an image to support the description.
end if
|  | select_one | yesno | leaks_electronic_compartment | 27. Are there any signs of leaks in the electronics compartment? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert(`Onsite:\nUse appropriate tools to access the electronics compartment. Inspect pipes, capillaries, and thermostat cable seal for leaks or moisture.\n\nRemote:\nSection B cannot be completed remotely. Ask an onsite user for this information.`)">[?]</span>
if | selected(data('leaks_electronic_compartment'), 'Yes')
|  | string |  | leaks_electronic_compartment_detail | 27.1. Describe the problem, location/component(s), and severity:
|  | image |  | leaks_electronic_compartment_image | 27.2. Upload an image to support the description.
end if
|  | select_one | yesno | fan_working_electronic_compartment | 28. Is the fan in the electronics compartment working? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Access the electronics compartment by removing the access cover/vent using appropriate tools. Observe whether the fan rotates when the compressor is running. If the compressor is not running, open the refrigerator door and wait 10 minutes to trigger it. Then recheck fan operation.\n\nRemote: Section B cannot be completed remotely. Please ask an onsite user to verify fan functionality.')">[?]</span>
|  | select_one | yesno | illicit_repairs_electronic_compartment | 29. Are there any indications of illicit repairs in the electronics compartment? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Access the electronics compartment by removing the access cover/vent using appropriate tools. Check for indications of illicit repairs—such as snipped wires, post-manufacturing additions, or incorrect components. Refer to a circuitry diagram if available to verify.\n\nRemote: Section B cannot be completed remotely. Ask a user who is onsite.')">[?]</span>
if | selected(data('illicit_repairs_electronic_compartment'), 'Yes')
|  | string |  | illicit_repairs_electronic_compartment_detail | 29.1. Describe the observation and the location/component(s) involved: <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Enter details about any repair(s) observed in the electronics compartment or other areas.\n\nRemote: Section B cannot be completed remotely. Ask a user who is on site for these details.')">[?]</span>
|  | image |  | illicit_repairs_electronic_compartment_image | 29.2. Upload an image to support the description.
end if
|  | decimal |  | operating_range_voltage_stabilizer | 30. Rated operating range of voltage stabilizer (volts)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Check the manufacturer label, other markings on the power quality device, and/or the device manual for the range of acceptable input voltage. If unknown, leave blank.\n\nRemote: Ask a user who is onsite. Section B cannot be completed remotely.')">[?]</span>
|  | decimal |  | input_voltage_reading | 31. What is the input voltage reading (volts)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('This question only applies to voltage stabilizers that display the input voltage they are currently receiving from the power source.\n\nOnsite: Record the input voltage displayed (in volts).\nRemote: Ask a user who is onsite. Section B cannot be completed remotely.')">[?]</span>
|  | decimal |  | output_voltage_reading | 32. What is the output voltage reading (volts)? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('This question only applies to voltage stabilizers that display the output voltage they are currently delivering to the refrigerator.\n\nOnsite: Record the output voltage displayed (in volts).\nRemote: Ask the user who is onsite. Section B cannot be completed remotely.')">[?]</span>
|  | string |  | replaced_model_same_device | 33. If the voltage stabilizer was replaced within the month, what is the make and model of the new device? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Find the make and model labeled on the voltage stabilizer.\n\nIf the voltage stabilizer was not replaced within the month, enter \"N/A\".\nIf it was replaced within the month but the make and model cannot be identified, enter \"unknown\".\n\nRemote: Ask a user who is onsite. Section B cannot be completed remotely.')">[?]</span>
|  | select_one | yes_no_unable | replaced_model_new_device | 34. If the voltage stabilizer was replaced within the month, is it the same make and model that was previously used with the refrigerator? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Reference inventory data collected at the start of PMM to determine if the new voltage stabilizer is of the same make and model as the one previously used with the refrigerator.\n\nRemote: Ask a user who is onsite. Section B cannot be completed remotely.')">[?]</span>
|  | select_one | provided_or_purchased | replaced_provided_or_purchased | 35. If the voltage stabilizer was not replaced within the month, was the voltage stabilizer provided with the refrigerator or purchased separately? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Ask the user if the voltage stabilizer was delivered with the refrigerator when it was first installed, or if it was purchased separately by the health facility or district.\n\nRemote: Ask a user who is onsite. Section B cannot be completed remotely.')">[?]</span>
|  | select_one | yesno | physical_damage_voltage_stabilizer | 36. Is there any visible physical damage to the voltage stabilizer? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Check the voltage stabilizer for any signs of damage, such as corrosion or dents.\n\nRemote: Ask the user to check the voltage stabilizer for visible damage.\n\nNote: Section B cannot be completed remotely without onsite verification.')">[?]</span>
|  | select_one_with_other | problems_with_connection | problems_connections_voltage_stabilizer | 37. Are there any problems with connections to the voltage stabilizer and/or the power cable? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Check the connections of the voltage stabilizer to the refrigerator power cable and power source. Look for loose connections, damaged cables, or signs of burning or dents on plugs/receptacles. Confirm if plug types are compatible.\n\nRemote: Ask the onsite user to perform these checks and report any issues.\n\nNote: Section B cannot be completed remotely.')">[?]</span> | [7 empty] | data('problems_with_connection') != ' ' | Other field is empty. Please specify.
|  | select_one | yesno | voltage_stabilizer_available_use | 38. If a voltage stabilizer for the refrigerator is available on site but is not currently in use, should this device be put into use with the refrigerator? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Ask if a voltage stabilizer was delivered with or purchased for this refrigerator. If one is available but not in use, ask why. Check for damage or plug incompatibility. If usable, connect and power it on (some stabilizers may delay ~5–10 mins).\n\nRemote: Section B cannot be completed remotely.')">[?]</span>
if | data('power_source') === 'solar'
|  | select_one | yesno | solar_system_grounded | 39. Is the solar system grounded? <span style="color:#007bff; border-bottom: 1px dotted #007bff; cursor: pointer;" onclick="alert('Onsite: Check if there is a cable connecting the solar array to a ground rod submerged in the earth within a ground pit. If such a connection exists, the system is grounded.\n\nRemote: Section B cannot be completed remotely. Please ask a user who is onsite.')">[?]</span>
end if
|  | select_one | yesno_section_b | failure_cause_known_b | 40. Has the cause of the failure been identified? | [5 empty] | TRUE
end screen
end if
if | selected(data('failure_cause_known_b'), 'Yes')
|  | linked_table | linked_failure_reporting |  | <div style="padding: 12px 16px; background-color: #fdecea; border-left: 5px solid #e74c3c; font-family: Arial, sans-serif; margin-bottom: 10px;"> <h2 style="margin: 0; color: #c0392b;">🚨 Failure Reporting</h2> </div>
end if
begin screen
|  | note |  | title_information_for_technician | <div style="padding: 12px 16px; background-color: #f1f9f4; border-left: 5px solid #27ae60; font-family: Arial, sans-serif; margin-bottom: 10px;"> <h2 style="margin: 0; color: #2e7d32;">📝Information for Technicians</h2> </div>
|  | select_one | yesno | tfa_visit_needed | Is Failure Analysis visit needed?
|  | string |  | repairs_concerns_technician | Enter any repairs attempted, or concerns for technician:
|  | string |  | replacement_parts_needed | Enter any replacement part(s) needed:
end screen
|  | note |  |  | Survey completed. Pressing NEXT will save and exit.

## model
type | name | isSessionVariable | elementType
string | refrigerator_id
string | model_id
string | manufacturer
string | facility_name
birthdate | followup_date
string | followup_uuid
string | power_source

## choices
choice_list_name | data_value | display.title.text | display.title.text.es | display.title.text.fr
yesno | Yes | Yes
yesno | No | No
refrigerator_place | Bottom of vaccine compartment | Bottom of vaccine compartment
refrigerator_place | Walls of vaccine compartment | Walls of vaccine compartment
refrigerator_place | On shelf or basket | On shelf or basket
refrigerator_place | Inside of door | Inside of door
yesnou | Yes | Yes
yesnou | No | No
yesnou | Unknown | Unknown
time | Weekly | Weekly
time | Monthly | Monthly
time | Yearly | Yearly
time | Never | Never
solar_array_location | Ground | Ground
solar_array_location | Roof | Roof
solar_array_location | Pole | Pole
main_component | cooling | Cooling system
main_component | electrical | Electrical system
main_component | electrical_solar | Electrical system (solar specific)
main_component | structural | Structural Components
main_component | not_applicable | Not applicable – no build quality issues reported
yes_no_unable | Yes | Yes
yes_no_unable | No | No
yes_no_unable | Unable to determine | Unable to determine
iceliner_intact | Yes, the ice liner is intact and in the correct place | Yes, the ice liner is intact and in the correct place
iceliner_intact | No, the ice liner is not intact and/or not in the correct place | No, the ice liner is not intact and/or not in the correct place
iceliner_intact | This refrigerator is not designed to operate with an ice liner | This refrigerator is not designed to operate with an ice liner
iceliner_intact | Unknown/no information | Unknown/no information
provided_or_purchased | Provided with refrigerator | Provided with refrigerator
provided_or_purchased | Purchased separately | Purchased separately
provided_or_purchased | Unable to determine | Unable to determine
problems_with_connection | Loose connection | Loose connection
problems_with_connection | Cable damage | Cable damage
problems_with_connection | Plug/Receptacle damage | Plug/Receptacle damage
problems_with_connection | Refrigerator plug type and voltage stabilizer receptacle type mismatched | Refrigerator plug type and voltage stabilizer receptacle type mismatched
onsiteremote | Onsite | Onsite
onsiteremote | Remote | Remote
yesno_section_a | Yes | Yes - continue to report the failure and causes
yesno_section_a | No | <span onclick="alert(`Failure cause not known. Please proceed to Section B.`)">No - continue on to Section B with more detailed follow-up questions</span>
yesno_section_b | Yes | Yes
yesno_section_b | No | <span onclick="alert(`Follow up has not identified the cause of failure. Please complete the Failure Analysis Survey.`)">No</span>
cooling_system | compressor | Compressor
cooling_system | piping | Piping
cooling_system | internal_condenser | Internal condenser
cooling_system | external_condenser | External condenser
cooling_system | filter_dryer | Filter dryer
cooling_system | capillary | Capillary
cooling_system | thermostat_sensor/external_display_sensor | Thermostat sensor/external display sensor
cooling_system | fan | Fan
cooling_system | compressor_controller/starting_device | Compressor controller/starting device
cooling_system | thermostat_controller | Thermostat controller
cooling_system | starting_capacitor | Starting capacitor
electrical_system | power_cable | Power cable(s)
electrical_system | keypad | Keypad
electrical_system | external_temperature_display | External temperature display
electrical_system | power_switch | Power switch
electrical_system | power_LED | Power LED
electrical_system | thermostat_sensor/external_display_sensor | Thermostat sensor/external display sensor
electrical_system | fan | Fan
electrical_system | compressor_controller/starting_device | Compressor controller/starting device
electrical_system | thermostat_controller | Thermostat controller
electrical_system | starting_capacitor | Starting capacitor
electrical_solar_system | solar_panels | Solar panels
electrical_solar_system | solar_array_support_structure/mounting | Solar array support structure/mounting
electrical_solar_system | solar_battery | Solar battery
electrical_solar_system | solar_wiring | Solar wiring
structural_components | casing | Casing
structural_components | base | Base
structural_components | access_cover | Access cover
structural_components | vents | Vents
structural_components | condensation_drainage | Condensation drainage
structural_components | handle | Handle
structural_components | gasket | Gasket
structural_components | alignment | Alignment
structural_components | hinges | Hinges
structural_components | latch/clasp | Latch/clasp
structural_components | lock/key | Lock/key
structural_components | ice_liner/water_packs | Ice liner/water packs
structural_components | shelving/baskets | Shelving/baskets
structural_components | vaccine_compartment_walls | Vaccine compartment walls
structural_components | vaccine_compartment_bottom | Vaccine compartment bottom
structural_components | grate | Grate
structural_components | drain | Drain
structural_components | sensor_mount_or_seal | Sensor mount or seal
structural_components | base_plate | Base plate

## queries
query_name | query_type | linked_form_id | linked_table_id | selection | selectionArgs | newRowInitialElementKeyToValueMap | openRowInitialElementKeyToValueMap
linked_failure_reporting | linked_table | failure_reporting | failure_reporting | followup_uuid = ? | [ data('followup_uuid') ] | {followup_uuid: data('followup_uuid')} | {}

## settings
setting_name | value | display.title.text | display.title.text.es | display.title.text.fr | display.locale.text | display.locale.text.es | display.locale.text.fr
form_id | follow_up
table_id | follow_up
survey |  | Follow-up
form_version | 20210930
instance_name | followup_date
english |  |  |  |  | English | Inglés | Anglais
es |  |  |  |  | Spanish | Español | Espagnol
fr |  |  |  |  | French | Francés | Français

## properties
partition | aspect | key | type | value
Table | default | defaultViewType | string | LIST
Table | default | detailViewFileName | string | config/tables/follow_up/html/follow_up_detail.html
Table | default | listViewFileName | string | config/tables/follow_up/html/follow_up_list.html
Table | security | unverifiedUserCanCreate | boolean | false
Table | security | defaultAccessOnCreation | string | HIDDEN
FormType | default | FormType.formType | string | SURVEY
SurveyUtil | default | SurveyUtil.formId | string | wrong_form